# 5 Whys Root Cause Worksheet

Attach to the investigation or CAPA record it supports.

How to use this worksheet Write one precise problem statement first. A vague problem gives vague answers. Ask why the previous answer happened, not who caused it. Stop when the answer is a cause you can act on that, once removed, prevents recurrence; that may take three whys or seven. Check every answer against evidence before asking the next why. If an answer has two possible causes, follow each in its own column or worksheet. "Human error" or "operator did not follow the SOP" is never the last answer: ask why the error was possible. A worked example is on page 2.

| Investigation / CAPA number |  | Date |  |
| --- | --- | --- | --- |
| Facilitator |  | Team |  |

Problem statement What, where, when and how much, compared with the requirement.

| Why? | Answer (cause) | Evidence that confirms it |
| --- | --- | --- |
| Why 1 |  |  |
| Why 2 |  |  |
| Why 3 |  |  |
| Why 4 |  |  |
| Why 5 |  |  |

| Root cause |  |
| --- | --- |
| Action that removes the root cause |  |
| Owner / due date / CAPA number |  |

## Worked example

Illustrative example. Delete this page before issuing the worksheet.

| Problem statement | On 2 November 2026 the pH meter in QC lab 2 was used for release testing of two batches while its calibration had expired on 28 October 2026. |
| --- | --- |

| Why? | Answer (cause) | Evidence that confirms it |
| --- | --- | --- |
| Why 1 | The analyst did not know the calibration had expired. | Interview; instrument had no status label. |
| Why 2 | Calibration status is only tracked in the lab supervisor’s spreadsheet, not on the instrument. | Spreadsheet reviewed; no label procedure. |
| Why 3 | The supervisor was on leave and nobody else reviewed the spreadsheet. | Leave schedule; no deputy named. |
| Why 4 | The calibration SOP names no deputy owner and requires no status check before use. | SOP-QC-005 v2, sections 4 and 5. |
| Why 5 | The SOP was written when the lab had one instrument and was not reassessed when the lab expanded. | SOP history; periodic review not done. |

| Root cause | SOP-QC-005 does not require a calibration status check before use or a deputy owner for the schedule. |
| --- | --- |
| Action that removes the root cause | Revise SOP-QC-005: status labels, pre-use check, deputy owner, shared schedule. Train analysts. Effectiveness: no use of out-of-calibration instruments in a logbook review after 3 months. |